## Read Me
Measurement Uncertainty Budget — Spectrophotometric Method (template)
Author: Mohamed Sayed Khammeri — Laboratory Manager, ISO/IEC 17025. Licence: CC BY 4.0.
ALL DATA IN THIS WORKBOOK ARE FICTIONAL. It contains no proprietary, client or accreditation-body information, and no text reproduced from any standard.
Purpose
Bottom-up (GUM / EURACHEM-CITAC Guide CG4) uncertainty estimate for a single-wavelength spectrophotometric determination with an external linear calibration. The worked example is orthophosphate (as P) in water by the molybdenum-blue reaction, after a 10 → 50 mL dilution.
Measurement model
c = x0 × Vf / Va where x0 = (ȳ0 − a) / b is read from the calibration line (Calibration sheet).
Relative standard uncertainties are combined in quadrature because the model is a product/quotient; repeatability and recovery enter as multiplicative factors equal to 1.
How to use
Yellow cells with blue text = inputs. Replace them with your own calibration, readings, glassware tolerances, certificate values, replicate and recovery data.
Green-shaded cells = key results. Every other number is a live formula — do not overwrite.
Green text = value pulled from another sheet.
Sheets: Calibration (regression and u(x0)) → Inputs (volumetric, reference standard, precision, recovery) → Budget (combination, contributions, expanded uncertainty, reported result).
Assumptions and limits (state these in your own procedure)
1. Calibration standards are assumed to carry negligible uncertainty on the x-axis compared with the absorbance scatter (ordinary least squares); the uncertainty of the stock standard is added separately.
2. Working-standard dilution steps are omitted for brevity. Add one row per pipette/flask used to prepare them in your real method.
3. The calibration term and the repeatability term partly overlap (both contain instrument noise). Keeping both is slightly conservative and is common practice; say so in your procedure.
4. Glassware: tolerance treated as a rectangular distribution (÷√3); temperature effect = V × ΔT × γ(water) ÷ √3.
5. Recovery is tested for significance (¦1 − R¦ / u(R) ≤ 2). If not significant, the result is not corrected and u(R) is included. If significant, correct the result or investigate — the Budget sheet flags this.
6. k = 2 gives ~95 % coverage assuming enough effective degrees of freedom. Check with Welch–Satterthwaite if a Type A term with few replicates dominates.
References (cite; do not copy)
JCGM 100:2008 (GUM) · EURACHEM/CITAC Guide CG4 'Quantifying Uncertainty in Analytical Measurement', 3rd ed. · ISO/IEC 17025:2017 §7.6 · Example method family: ISO 6878 (phosphorus, ammonium molybdate spectrometric method).

## Budget
Uncertainty budget and reported result
Model: c = x0 × Vf / Va (× repeatability factor × recovery factor, both = 1)
Quantity | Value | Unit
x0 — concentration in measuring solution | 0.51417 | mg P/L
Dilution factor Vf / Va | 5 | -
Result c | 2.5708 | mg P/L
Uncertainty component | Value x | Standard uncertainty u(x) | Unit | Type / distribution | Relative u(x)/x | Contribution to u_c² (%)
Calibration curve (x0) | 0.51417 | 0.00088528 | mg P/L | A — regression (normal) | 0.1722% | 1.357%
Stock standard (certificate) | 1000 | 1 | mg P/L | B — certificate U/k (normal) | 0.1% | 0.4578%
Sample aliquot Va (pipette) | 10 | 0.012524 | mL | B — tolerance + temperature (rectangular) | 0.1252% | 0.718%
Final volume Vf (flask) | 50 | 0.042285 | mL | B — tolerance + temperature (rectangular) | 0.08457% | 0.3274%
Repeatability / intermediate precision | 2.545 | 0.033747 | mg P/L | A — replicate QC results | 1.326% | 80.49%
Recovery R (method bias) | 0.99467 | 0.0059981 | - | A — spike recovery, s/√n | 0.603% | 16.65%
Total | [5 empty] | 100%
Combined and expanded uncertainty | Value | Unit
Combined relative standard uncertainty u_c(c)/c | 1.478% | -
Combined standard uncertainty u_c(c) | 0.037997 | mg P/L
Coverage factor k | 2 | -
Expanded uncertainty U = k · u_c | 0.075994 | mg P/L
Relative expanded uncertainty U/c | 2.956% | -
Reported result | c = 2.57 ± 0.08 mg P/L (k = 2, approx. 95 % coverage)
Recovery significance test | Value | Unit
t = ¦1 − R¦ / u(R) | 0.88916 | -
Conclusion | Bias not significant (t <= 2): result not corrected, u(R) included
Dominant contribution | Component
Largest share of u_c² | Repeatability / intermediate precision
Improve the method where the largest share sits — a smaller term barely changes u_c because terms add in quadrature.

## Calibration
Calibration line and uncertainty of x0
Ordinary least-squares regression: absorbance (y) vs concentration (x). Fictional data.
Level | Conc. x (mg P/L) | Absorbance y | Predicted ŷ | Residual y − ŷ | Regression statistic | Value
Std 0 | 0 | 0.003 | 0.0032304 | -0.00023038 | Slope b (AU per mg P/L) | 0.51787
Std 1 | 0.1 | 0.055 | 0.055017 | -1.7065e-05 | Intercept a (AU) | 0.0032304
Std 2 | 0.2 | 0.107 | 0.1068 | 0.00019625 | Coefficient of determination R² | 0.99999
Std 3 | 0.4 | 0.211 | 0.21038 | 0.00062287 | Residual standard deviation s(y/x) | 0.00056911
Std 4 | 0.6 | 0.313 | 0.31395 | -0.00095051 | Number of calibration points n | 7
Std 5 | 0.8 | 0.418 | 0.41752 | 0.00047611 | Mean concentration x̄ | 0.44286
Std 6 | 1 | 0.521 | 0.5211 | -9.727e-05 | Mean absorbance ȳ | 0.23257
[5 empty] | Sxx = Σ(xi − x̄)² | 0.83714
Sample reading | Absorbance y0 |  |  |  | Sample result | Value
Reading 1 | 0.268 |  |  |  | Number of sample readings p | 2
Reading 2 | 0.271 |  |  |  | Mean sample absorbance ȳ0 | 0.2695
Reading 3 |  |  |  |  | x0 = (ȳ0 − a) / b (mg P/L, measuring solution) | 0.51417
Reading 4 |  |  |  |  | u(x0) = s(y/x)/b · √(1/p + 1/n + (ȳ0 − ȳ)² / (b²·Sxx)) | 0.00088528
Reading 5 |  |  |  |  | Relative u(x0) / x0 | 0.1722%
Leave unused readings empty (up to 5). |  |  |  |  | Range check | Within calibration range

## Inputs
Inputs — volumetric, reference standard, precision, recovery
Yellow = your data. Source column documents where each value comes from.
Parameter | Value | Unit | Source / note
Volumetric operations
Sample aliquot Va (pipette) | 10 | mL | Class A volumetric pipette, nominal volume
Pipette tolerance (±) | 0.02 | mL | Manufacturer / class A tolerance (fictional)
Final volume Vf (flask) | 50 | mL | Class A volumetric flask, nominal volume
Flask tolerance (±) | 0.06 | mL | Manufacturer / class A tolerance (fictional)
Laboratory temperature variation (±) | 4 | °C | From lab environmental records (fictional)
Volume expansion coefficient of water γ | 0.00021 | /°C | Physical constant, ~20 °C
Reference standard (stock solution)
Stock standard concentration | 1000 | mg P/L | Certificate of the CRM / stock standard (fictional)
Expanded uncertainty on certificate U | 2 | mg P/L | Certificate (fictional)
Coverage factor on certificate k | 2 | - | Certificate
Precision — replicate results of a QC sample (intermediate conditions)
QC replicate 1 | 2.51 | mg P/L | Different days/analysts if possible
QC replicate 2 | 2.58 | mg P/L
QC replicate 3 | 2.55 | mg P/L
QC replicate 4 | 2.49 | mg P/L
QC replicate 5 | 2.6 | mg P/L
QC replicate 6 | 2.54 | mg P/L
QC replicate 7 | 2.57 | mg P/L
QC replicate 8 | 2.52 | mg P/L
QC replicate 9 | 2.56 | mg P/L
QC replicate 10 | 2.53 | mg P/L
Recovery — spiked samples (method bias)
Spike recovery 1 | 98.2 | % | (found − unspiked) / added × 100
Spike recovery 2 | 101.5 | %
Spike recovery 3 | 99.4 | %
Spike recovery 4 | 97.8 | %
Spike recovery 5 | 100.9 | %
Spike recovery 6 | 99 | %
Reporting
Coverage factor k for the result | 2 | - | ~95 % coverage
Number of independent results averaged in the reported value | 1 | - | 1 = single determination