## Read Me
Method Validation Workbook — Quantitative Chemical Method (template)
Author: Mohamed Sayed Khammeri — Laboratory Manager, ISO/IEC 17025. Licence: CC BY 4.0.
ALL DATA IN THIS WORKBOOK ARE FICTIONAL. It contains no proprietary, client or accreditation-body information, and no text reproduced from any standard.
Purpose
Collects the performance characteristics usually required to show a quantitative method is fit for purpose (ISO/IEC 17025:2017 §7.2.2), and compares each one with a pre-defined acceptance criterion. Example: orthophosphate (as P) in water by molybdenum-blue spectrophotometry — the same fictional method as the uncertainty workbook.
Sheets
Summary — every characteristic, its criterion (editable) and a PASS / FAIL verdict. This is the page an assessor reads first.
Linearity — calibration regression, residuals (absolute and %), R², residual plot.
LOD-LOQ — from replicate low-level / blank determinations: LOD = 3·s0, LOQ = 10·s0 (EURACHEM approach), expressed in the measuring solution and in the sample.
Precision — one-way ANOVA on a QC sample (p days × n replicates): repeatability s_r, between-day s_b, intermediate precision s_I, repeatability limit r = 2.8·s_r.
Trueness — CRM bias test (¦bias¦ compared with U of the difference) and spike recovery at three levels.
How to use
Yellow cells with blue text = inputs (your data and your acceptance criteria). Criteria must be fixed in the validation plan BEFORE the experiments.
Green-shaded cells = key results. Every other number is a live formula.
Green text = value pulled from another sheet.
Limits of this template
1. Linearity is judged on R² and residuals only. For a stricter check add replicate standards and a lack-of-fit F-test or Mandel's test.
2. LOD/LOQ from 10 independent low-level determinations; for methods with a non-zero blank, use blank-corrected results and state it.
3. Precision uses a balanced design (same n per day). Unbalanced designs need a different ANOVA formula.
4. Selectivity/interferences, robustness and measuring range are method-specific and should be documented in the validation report alongside this workbook.
References (cite; do not copy)
EURACHEM Guide 'The Fitness for Purpose of Analytical Methods', 2nd ed. (2014) · ISO/IEC 17025:2017 §7.2 · ISO 5725-2 (precision, ANOVA) · ERM Application Note 1 (comparing a result with a certified value).

## Summary
Validation summary — orthophosphate (as P) in water, molybdenum-blue spectrophotometry
Fictional example. Criteria are set in the validation plan before the experiments.
Characteristic | Result | Criterion | Rule | Verdict
Linearity — R² | 0.99997 | 0.995 | Result ≥ criterion | PASS
Linearity — max ¦residual %¦ | 1.352% | 5% | Result ≤ criterion | PASS
LOQ in sample (mg P/L) | 0.054006 | 0.1 | Result ≤ required LOQ | PASS
Repeatability RSD_r | 0.6825% | 2% | Result ≤ criterion | PASS
Intermediate precision RSD_I | 1.49% | 3% | Result ≤ criterion | PASS
CRM relative bias | -0.1333% | 3% | ¦Result¦ ≤ criterion | PASS
Spike recovery — minimum | 98% | 90% | Result ≥ criterion | PASS
Spike recovery — maximum | 102% | 110% | Result ≤ criterion | PASS
CRM bias significance test | No (¦bias¦ <= U(bias)) | not significant | ¦bias¦ ≤ U(bias) | PASS
Overall conclusion | All criteria met - method fit for the intended purpose
Also document in the validation report: selectivity / interferences, working range, robustness, measurement uncertainty (see the uncertainty workbook), and the decision of the technical manager.

## Linearity
Linearity — calibration regression and residuals
Fictional data. Blank level excluded from % residual.
Level | Conc. x (mg P/L) | Absorbance y | Predicted ŷ | Residual y − ŷ | Residual % | Statistic | Value
Std 0 | 0 | 0.004 | 0.0042952 | -0.00029522 |  | Slope b | 0.51578
Std 1 | 0.1 | 0.056 | 0.055874 | 0.00012628 | 0.226% | Intercept a | 0.0042952
Std 2 | 0.2 | 0.106 | 0.10745 | -0.0014522 | -1.352% | R² | 0.99997
Std 3 | 0.4 | 0.212 | 0.21061 | 0.0013908 | 0.6604% | s(y/x) | 0.0011697
Std 4 | 0.6 | 0.315 | 0.31377 | 0.0012338 | 0.3932% | Max ¦residual %¦ (x > 0) | 1.352%
Std 5 | 0.8 | 0.417 | 0.41692 | 7.6792e-05 | 0.01842% | Intercept as % of mid-range signal | 1.88%
Std 6 | 1 | 0.519 | 0.52008 | -0.0010802 | -0.2077% | Indicative LOD from curve = 3.3·s(y/x)/b | 0.0074839

## LOD-LOQ
Limit of detection and limit of quantification
10 independent determinations of a blank / very-low-level sample, each through the full procedure. Fictional data.
Determination | Result (mg P/L, measuring solution) | Statistic | Value | Unit
Blank 1 | 0.003 | Dilution factor applied to samples | 5 | -
Blank 2 | 0.005 | Number of determinations | 10 | -
Blank 3 | 0.002 | Mean | 0.0035 | mg P/L
Blank 4 | 0.004 | Standard deviation s0 | 0.0010801 | mg P/L
Blank 5 | 0.004 | LOD = 3·s0 (measuring solution) | 0.0032404 | mg P/L
Blank 6 | 0.003 | LOQ = 10·s0 (measuring solution) | 0.010801 | mg P/L
Blank 7 | 0.005 | LOD in sample | 0.016202 | mg P/L
Blank 8 | 0.002 | LOQ in sample | 0.054006 | mg P/L
Blank 9 | 0.004
Blank 10 | 0.003 | Check that the LOQ is below the lowest calibration standard used for reporting, and below the lowest limit your clients need to verify.

## Precision
Precision — repeatability and intermediate precision (one-way ANOVA)
QC sample analysed in triplicate on 5 different days (different analyst / reagent batch where possible). Fictional data.
Day (group) | Rep 1 | Rep 2 | Rep 3 | Day mean | Σ(x − day mean)²
Day 1 | 2.54 | 2.57 | 2.55 | 2.5533 | 0.00046667
Day 2 | 2.6 | 2.58 | 2.61 | 2.5967 | 0.00046667
Day 3 | 2.51 | 2.53 | 2.5 | 2.5133 | 0.00046667
Day 4 | 2.56 | 2.55 | 2.59 | 2.5667 | 0.00086667
Day 5 | 2.58 | 2.62 | 2.6 | 2.6 | 0.0008
ANOVA statistic | Value | Unit
p — number of days | 5 | -
n — replicates per day | 3 | -
Grand mean | 2.566 | mg P/L
MS within (repeatability) | 0.00030667 | (mg P/L)²
MS between (days) | 0.0037733 | (mg P/L)²
s_r — repeatability SD | 0.017512 | mg P/L
s_b — between-day SD | 0.033993 | mg P/L
s_I — intermediate precision SD | 0.038239 | mg P/L
RSD_r | 0.6825% | -
RSD_I | 1.49% | -
Repeatability limit r = 2.8·s_r | 0.049033 | mg P/L
Intermediate precision limit = 2.8·s_I | 0.10707 | mg P/L
If MS between < MS within, the between-day variance is set to 0 (s_I = s_r).

## Trueness
Trueness — CRM bias and spike recovery
Fictional data.
A. Certified reference material
Certified value | 2.5 | mg P/L
Expanded uncertainty on certificate U | 0.05 | mg P/L
Coverage factor on certificate k | 2 | -
CRM replicate | Result (mg P/L) |  | Statistic | Value | Unit
Replicate 1 | 2.47 |  | n | 6 | -
Replicate 2 | 2.53 |  | Mean | 2.4967 | mg P/L
Replicate 3 | 2.49 |  | s | 0.028048 | mg P/L
Replicate 4 | 2.52 |  | Bias Δ = mean − certified | -0.0033333 | mg P/L
Replicate 5 | 2.46 |  | Relative bias | -0.1333% | -
Replicate 6 | 2.51 |  | Recovery | 99.87% | -
|  |  | u(Δ) = √(s²/n + (U_CRM/k)²) | 0.027497 | mg P/L
|  |  | U(Δ) = 2·u(Δ) | 0.054995 | mg P/L
|  |  | Bias significant? | No (¦bias¦ <= U(bias))
B. Spike recovery at three levels (duplicates)
Level | Unspiked result (mg P/L) | Spike added (mg P/L) | Spiked result (mg P/L) | Recovery %
Low | 0.52 | 0.5 | 1.01 | 98%
Low | 0.52 | 0.5 | 1.03 | 102%
Mid | 0.52 | 2 | 2.49 | 98.5%
Mid | 0.52 | 2 | 2.54 | 101%
High | 0.52 | 4 | 4.47 | 98.75%
High | 0.52 | 4 | 4.55 | 100.8%
|  |  | Mean recovery | 99.83%
|  |  | Minimum | 98%
|  |  | Maximum | 102%